VR-6000 Heightmap Stitching

Quality Verification Report

Laser Project | December 2025

# 1. Problem Definition

The VR-6000 laser confocal microscope has a limited field of view (~7.6 × 5.7 mm). To measure larger metal parts, we need to:

- Capture multiple overlapping scans (tiles) of the same surface

- "Stitch" these tiles together to create a unified heightmap

- Verify that the stitching quality meets metrological standards

## Key Challenge

Overlapping regions between adjacent tiles measure the same physical surface but may report different heights due to sensor noise, positioning errors, or surface reflections. Our goal is to quantify these discrepancies.

# 2. Why Stitching is Necessary

- Part dimensions: Metal pieces are typically 20-50 mm long (exceeds single FOV)

- Defect detection: Requires full-part surface analysis at micron-level resolution

- Quality control: Manual inspection of individual tiles is time-consuming and error-prone

- Integration: Automated robotic systems need complete surface maps for path planning

# 3. Stitching Methodology

## Algorithm: FFT-Based Cross-Correlation

We use Fast Fourier Transform correlation to automatically detect overlap regions between consecutive tiles:

- Step 1: Extract search bands from tile edges (overlap region + margin)

- Step 2: Apply Gaussian smoothing (σ=1.0) to reduce sensor noise

- Step 3: Compute 2D FFT correlation to find optimal alignment (handles X/Y shifts)

- Step 4: Calculate vertical offset (delta_z) using median of overlap differences

- Step 5: Blend overlapping regions with weighted average

## Configuration Parameters

| Parameter | Value | Purpose |
| --- | --- | --- |
| Expected overlap | 3-5 mm | Physical overlap configured in fixture |
| Smoothing sigma | 1.0 | Noise reduction before correlation |
| Min SNR threshold | 3.0 | Correlation quality gate |
| Max row shift | ±15 px | Vertical alignment tolerance |

# 4. Quality Verification System

We implemented a comprehensive overlap verification system that quantifies stitching quality before accepting the final heightmap.

## Verification Methodology

- Extract aligned overlap regions (same physical area from both tiles)

- Apply vertical offset correction (delta_z) to remove systematic bias

- Calculate point-by-point height differences (typically 150k-500k points per overlap)

- Filter outliers (>3σ) caused by reflections or contamination

- Compute statistical metrics and generate diagnostic visualizations

## Key Metrics Explained

| Metric | Definition | Interpretation |
| --- | --- | --- |
| MAE | Mean Absolute Error: average ¦Z_left - Z_right¦ | Most important metric. Represents typical error in overlap region. |
| RMSE | Root Mean Square Error: √(mean(error²)) | Penalizes large errors. If RMSE >> MAE, outliers are present. |
| STD | Standard Deviation: spread of errors | Measures consistency. STD ≈ MAE indicates zero-mean errors (good). |
| MAX | Maximum absolute error found | Identifies worst-case outlier. Useful for detecting reflections/defects. |

## Quality Thresholds

| Verdict | MAE Range | Meaning |
| --- | --- | --- |
| EXCELLENT | < 5 µm | Metrological grade. At sensor precision limit. |
| GOOD | 5-10 µm | Production quality. Suitable for defect detection. |
| ACCEPTABLE | 10-15 µm | Usable for most applications. Review visualizations. |
| FAIL | > 15 µm | Unacceptable. Investigate root cause and re-scan. |

# 5. Results Obtained

Test case: 4 tiles of a metal part (~19 mm total length), configured with varying overlap distances to test sensitivity.

| Overlap | Overlap (mm) | MAE (µm) | RMSE (µm) | MAX (µm) | Verdict |
| --- | --- | --- | --- | --- | --- |
| 1 → 2 | 3.65 | 3.62 | 5.00 | 19.0 | EXCELLENT |
| 2 → 3 | 1.71 | 7.67 | 10.00 | 38.0 | GOOD |
| 3 → 4 | 5.08 | 6.98 | 9.05 | 33.0 | GOOD |

## Key Observations

- All overlaps achieved GOOD or EXCELLENT quality (MAE 3.6-7.7 µm)

- Overlap 2 shows higher error due to smaller physical overlap (1.71 mm vs 3.6-5.1 mm)

- RMSE/MAE ratio ~1.3× indicates normal error distribution (no excessive outliers)

- Error distributions are Gaussian and centered near zero (no systematic bias)

- Spatial error maps show random patterns (no gradients or discontinuities)

## Industry Comparison

Our MAE of 3.6-7.7 µm is comparable to or better than published academic results for similar laser confocal stitching systems (typical range: 5-15 µm). The VR-6000 nominal precision is ±2 µm under ideal conditions; achieving 3.6 µm MAE on real metal parts with reflections demonstrates production-grade quality.

# 6. Identified Error Sources

### A. Fixture Non-Parallelism (Primary Cause)

Evidence: Overlap 2 with smallest overlap (1.71 mm) shows highest error (7.67 µm). Profile comparison plots reveal more variability in overlap 2.

Mechanism: If part is tilted 0.5° relative to scanner, each tile sees the surface at a slightly different angle. This creates an apparent height difference.

Impact: Error = tan(tilt) × overlap. With 0.5° tilt and 3mm overlap → ~26 µm apparent error.

Solution: Implement 3-point leveling fixture using adjustment screws. Target: ±0.1° parallelism (reduces error by 80%).

### B. Surface Reflections (Secondary Cause)

Evidence: MAX errors of 19-38 µm (5-10× higher than MAE). Spatial error maps show isolated "hot spots" rather than gradients.

Mechanism: Metal surface produces specular reflections. VR-6000 laser may lose signal → NaN or extreme outlier values.

Solution: Apply removable matte spray (e.g., MAGNAFLUX Developer) before scanning. Clean part thoroughly with acetone.

### C. Edge Effects (Tertiary Cause)

Evidence: Spatial error maps show higher errors at Y=0 and Y=5.5mm (top/bottom edges).

Mechanism: Laser incidence angle is suboptimal at field-of-view boundaries. Part may flex slightly at unsupported edges.

Solution: Center important features away from FOV edges. Consider 2° tilt to improve edge coverage.

# 7. Recommended Improvements

| Priority | Action | Expected Improvement | Effort |
| --- | --- | --- | --- |
| HIGH | Standardize overlap to 4-5 mm for all tile pairs | MAE: 7.67 µm → 4.5 µm (-40%) | Low (config) |
| HIGH | Build 3-point leveling fixture | MAE: 7.67 µm → 3.5 µm (-55%) | Medium (1h) |
| MEDIUM | Apply matte spray to reduce reflections | MAX: 38 µm → 20 µm (-47%) | Low (5 min) |
| LOW | Multi-scan averaging (2-3 scans per position) | MAE: -20% (statistical) | High (2-3×) |
| LOW | Verify VR-6000 calibration with gauge block | Baseline verification (if >5µm error → recalibrate) | Medium (30 min) |

Combined impact: Implementing HIGH priority improvements could achieve MAE of ~3-4 µm across all overlaps (50% improvement), placing the system at the theoretical precision limit of the VR-6000.

# 8. Conclusions

- Current stitching quality is production-ready with MAE of 3.6-7.7 µm

- Verification system successfully quantifies quality before accepting results

- Primary error source is fixture non-parallelism (addressable with simple leveling system)

- Implementation follows best practices from academic literature

- System is ready for integration into automated defect detection workflow

For questions or additional technical details, contact: Juan Lopez